## Sheet1
CH0 — ENV + Sampling Control
Channel | Functionality | Requirement | Why Do We Need It? | What Data Do We Get From It? | Industry-Grade Component | Low-Cost Component
CH0 | Temp + RH (Chamber) | Measure T/RH and compute dew point/AH for normalization and env_in_range gating. | VOC sensor response strongly depends on humidity and temperature; required for normalization and repeatable measurements | Temperature (°C), Relative Humidity (%RH), computed Dew Point (°C), Absolute Humidity (g/m³) | Sensirion SHT45 / SHT85 | Sensirion SHT41
CH0 | Pressure | Monitor chamber pressure to detect leak or abnormal airflow conditions. | Detect enclosure leaks, abnormal airflow behavior, or pressure instability affecting concentration | Absolute pressure (Pa or hPa), pressure trend over time | Bosch BMP390 | Bosch BMP388
CH0 | Near-Sample Temperature | Measure localized thermal condition to confirm stabilization before exposure. | Ensure sample reaches thermal equilibrium before VOC exposure; prevents thermal bias in readings | Local temperature near sample (°C), stabilization status (delta T over time) | PT100/PT1000 + MAX31865 | 10k NTC Thermistor + MCU ADC
CH0 | Flow Integrity (Inline) | Validate commanded airflow and generate flow_ok status flag. | Ensure correct airflow during baseline and exposure phases; detect fan degradation | Mass flow rate (L/min or slm), flow stability over time | Sensirion SFM3300/SFM3019 | Omron D6F-PH (low-range MEMS)
CH0 | DP Across Filter | Detect clog_risk and airflow restriction using differential pressure. | Detect gradual filter clogging and airflow restriction that changes VOC concentration | Differential pressure (Pa), clog trend (ΔP increase over time) | Honeywell TruStability HSC/SSC | NXP MPXV7002DP
CH0 | Fan RPM Feedback | Provide fan_ok health signal using tach feedback. | Detect fan stall, slowdown, or mechanical failure | Fan RPM (rotations per minute), pulse frequency | Delta / Sunon Industrial 4-wire PWM Fan | Generic 4-wire PWM Fan
CH0 | Current Monitoring | Detect stall/overload via actuator current signature. | Detect abnormal current draw indicating mechanical blockage or electrical fault | Current (A or mA), voltage (V), power (W) | TI INA226 | TI INA219
CH0 | Sampling Chamber Control | Ensure repeatable mixing, exposure, and recovery phases. | Maintain consistent airflow profile across SOP phases for repeatability | PWM duty cycle (%), control loop output, airflow stability metrics | Closed-loop PID airflow controller | MCU PWM control
CH0 | Run Gating Outputs | Provide warmup_done, baseline_stable, env_in_range, integrity flags to CH1–CH6. | Prevent CH1–CH6 from running when system conditions are unstable | Boolean status flags: warmup_done, baseline_stable, env_in_range, flow_ok, integrity_ok | Dedicated MCU + watchdog | MCU firmware flags
CH1 — VOC Fingerprint / MOX e-Nose
Channel | Functionality | Requirement | Why Do We Need It? | What Data Do We Get From It? | Industry-Grade Component | Low-Cost Component
CH1 | Sampling & Measurement | Execute controlled SOP cycle (warm-up → baseline → exposure → recovery). | Ensures repeatable concentration dynamics and stable baseline before exposure | Time-stamped phase markers, baseline resistance values, exposure response curves | STM32 + RTOS control | Basic Arduino framework
CH1 | MOX/MOS Array | Capture multi-dimensional VOC response (replaceable modular array). | Different MOX sensors respond differently to VOC mixtures; creates chemical fingerprint | Raw sensor resistance (Rs), ΔR/R0, conductance changes across multiple sensors | Bosch BME688 (x2) + Sensirion SGP40 | MQ2/MQ135 series
CH1 | Sensor Heat Control | Maintain stable heater profile with PWM/closed-loop control. | MOX sensitivity depends on heater temperature; unstable heating causes drift | Heater temperature setpoint, PWM duty cycle, heater current, temperature stability metrics | Closed-loop heater driver IC | PWM via MCU
CH1 | Controlled Airflow | Maintain repeatable concentration dynamics via CH0 airflow. | Exposure concentration and recovery speed depend on airflow consistency | Flow-aligned gas response curve, exposure rise rate, recovery slope | Sensirion SFM series | Fixed-speed fan
CH1 | Time-Series Sync | Align gas timestamps with NIR (CH4) and Vision (CH3). | Enables multi-modal fusion and consistent feature alignment | Synchronized timestamps across gas, vision, and spectral channels | RTC-synced hardware timestamp | Software timestamp
CH1 | Low-Noise Power + EMC | Provide stable ADC baseline and minimize electrical noise. | MOX signals are small analog changes; electrical noise corrupts features | Stable ADC counts, reduced baseline fluctuation (noise floor data) | LT3042 LDO + LC filters | AMS1117 + basic decoupling
CH1 | Drift & Repeatability | Track baseline shift and enforce compensation strategy. | MOX sensors drift over time; compensation required for long-term reliability | Baseline resistance (R0), drift trend over runs, compensation factors | Auto-baseline correction firmware | Manual baseline offset
CH1 | Continuous / On-Demand Mode | Support real-time monitoring and triggered batch measurement. | Allows both production-line monitoring and batch lab testing | Streaming sensor time-series or batch measurement datasets | RTOS scheduling | State-machine firmware
CH1 | Feature Extraction | Compute Δ, AUC, rise time, recovery slope features. | Raw time-series must be converted into discriminative features | ΔR, normalized response, AUC, rise time (t90), recovery constant, steady-state value | Edge AI processor (Orin NX) | Raspberry Pi 5
CH1 | Fingerprint Vector | Generate fixed-length embedding (e.g., 64-dim schema versioned). | Standardized representation for ML comparison and scoring | Fixed-length feature vector (e.g., 64-dim embedding), schema version ID | Embedded ML runtime | Lightweight ML library
CH1 | Scoring Engine | Output freshness, anomaly, adulteration index with confidence. | Converts fingerprint into actionable quality metrics | Freshness score, anomaly score, adulteration risk index, confidence level | TensorRT / Edge AI | Python ML inference
CH2 — Context Gases / Storage Atmosphere
Channel | Functionality | Requirement | Why Do We Need It? | What Data Do We Get From It? | Industry-Grade Component | Low-Cost Component
CH2 | Gas Coverage | Measure CO₂, O₂, H₂S, NH₃, SO₂, NO₂, CO as contextual markers. | Provides confirmatory and contextual markers to reduce false positives from MOX array | Gas concentration values (ppm or %), per-gas marker presence | Senseair S8 (CO₂), Alphasense O₂-A2, Alphasense toxic gas sensors | MH-Z19B (CO₂), Winsen ME2 series
CH2 | Measurement Modes | Support continuous, spot, and batch modes. | Enables flexibility for real-time monitoring and batch-based testing | Time-series concentration data (continuous) or discrete sampled readings (spot/batch) | Industrial NDIR/EC modules with streaming output | UART-based gas modules
CH2 | Controlled Cycle | Execute sample → stabilize → report → repeat cycle. | Ensures sensor stabilization before reporting; improves repeatability and accuracy | Stabilization time, steady-state concentration value, response time | Sensors with built-in stabilization control | Firmware-timed stabilization
CH2 | Timestamped Outputs | Output concentration, thresholds, context flags with timestamps. | Required for synchronization with CH1, CH3, CH4 and for audit traceability | Timestamped gas concentration (ppm), threshold flags, context status flags | RTC-synced industrial modules | MCU timestamping
CH2 | Cross-Channel Fusion | Fuse CO₂/RH/Temp data to reduce false alarms. | Environmental factors affect gas readings; fusion improves reliability | Compensated concentration values, fusion-adjusted confidence flags | Edge processor fusion engine | MCU rule-based fusion
CH2 | Calibration Management | Support recalibration policy and drift tracking. | Electrochemical/NDIR sensors drift over time; recalibration ensures accuracy | Zero offset, span correction factors, calibration timestamp, drift trend data | Auto-calibrating NDIR (ABC logic), calibration EEPROM | Manual firmware-based calibration
CH2 | Traceability Logs | Store raw + calibrated values with metadata. | Required for regulatory compliance and evidence tracking | Raw sensor output (mV or ADC counts), calibrated ppm values, sensor ID, run metadata | Secure signed logging system | Standard timestamped logging
CH2 | Low-Noise ADC | Ensure ppm-level resolution and accuracy. | Gas concentration changes may be small; requires high-resolution conversion | High-resolution ADC counts, filtered signal values, noise floor metrics | 24-bit ADC (ADS1220 class) | 16-bit ADC (ADS1115 class)
CH2 | Detection Specs | Meet defined range, response time, accuracy, and drift criteria. | Ensures system meets safety, food quality, or industrial detection standards | Measured gas range (ppm), response time (t90), accuracy %, long-term drift rate | Industrial-grade NDIR/EC sensors | Consumer-grade gas modules
CH2b — Electrochemical Selective Gases
Channel | Functionality | Requirement | Why Do We Need It? | What Data Do We Get From It? | Industry-Grade Component | Low-Cost Component
CH2b | EC Sensor Array | Provide ppm-level selective toxic/spoilage gas detection. | Provides highly selective and low-level detection of specific toxic or spoilage gases that MOX cannot reliably isolate | Raw sensor current (µA or nA), gas-specific response signal | Alphasense / CityTech EC sensors (NH₃, H₂S, SO₂, NO₂, CO) | Winsen ZE series EC sensors
CH2b | Analog Front End | Convert sensor current to stable digital signal (TIA + ADC). | EC sensors output very small currents; requires precise current-to-voltage conversion and high-resolution ADC | Converted voltage signal (mV), high-resolution ADC counts | TI LMP91000 + ADS1220 (24-bit ADC) | Discrete TIA (OPA333) + ADS1115
CH2b | Calibration Manager | Maintain zero/span, expiry gating, and calibration lineage. | EC sensors drift and expire; calibration ensures measurement validity and compliance | Zero offset value, span factor, calibration timestamp, expiry status | On-board calibration EEPROM + secure MCU | Firmware-based calibration storage
CH2b | Signal Conditioning | Apply filtering and cross-sensitivity guards. | Reduces electrical noise and compensates for temperature/humidity cross-effects | Filtered signal, temperature-compensated ppm estimate, stability metrics | Hardware + digital filtering with temperature compensation | Digital moving-average filtering in MCU
CH2b | Marker Extraction | Generate confidence-scored ppm estimates. | Converts conditioned signal into interpretable gas concentration with reliability score | Gas concentration (ppm), confidence score, threshold exceed flag | Model-based concentration estimation | Linear ppm conversion algorithm
CH2b | Fusion Output | Provide confirmatory markers to CH1/CH2 fusion. | Confirms or rejects anomalies detected by MOX; reduces false alarms | Confirmatory gas markers (ppm), fusion confidence flag | Multi-channel fusion engine (Edge processor) | MCU-based rule engine
CH2b | Platform Logs | Store calibrated values and metadata in evidence pack. | Required for auditability, regulatory traceability, and performance tracking | Raw current, calibrated ppm value, sensor ID, calibration metadata, timestamps | Secure signed logs with device ID | Standard timestamped logs
CH2b | Alerts | Trigger safety/quality alert signals with reason codes. | Enables immediate action for toxic gas or spoilage detection | Alert status (Warn/Alert), gas type, threshold level crossed, reason code | Configurable threshold engine with policy control | Fixed firmware threshold alerts
CH3 — Vision / AIV
Channel | Functionality | Requirement | Why Do We Need It? | What Data Do We Get From It? | Industry-Grade Component | Low-Cost Component
CH3 | Camera | Capture controlled RGB image for grading/defect analysis. | Capture high-quality, distortion-free image for grading and defect detection | RGB image (RAW/JPEG), resolution, exposure metadata | Industrial USB global shutter camera (e.g., Basler/Daheng) | Raspberry Pi Camera Module 3
CH3 | Lens / FOV | Maintain fixed geometry and calibrated field coverage. | Ensures consistent size measurement and spatial calibration | Field of view (mm coverage), pixel-to-mm scale factor | C-mount fixed-focus industrial lens | Fixed-focus M12 lens
CH3 | Colorimeter | Provide calibrated color reference for grading consistency. | Maintain color grading consistency across time and lighting changes | Reference color values (RGB/XYZ), ΔE color difference | Konica Minolta handheld colorimeter (reference calibration) | On-board color sensor (e.g., TCS34725)
CH3 | Lighting | Deliver uniform illumination under controlled enclosure. | Prevent shadows and brightness variation affecting grading | Illumination intensity (lux), uniformity metrics | Industrial ring/area light with diffuser | LED ring light module
CH3 | LED Driver | Maintain constant current with low flicker. | Ensures stable brightness and prevents flicker artifacts | LED current (mA), duty cycle %, brightness stability | Constant-current industrial LED driver | PWM-controlled LED driver
CH3 | Enclosure/Hood | Block ambient light interference. | Eliminates ambient light variation for consistent imaging | Ambient light suppression status | Light-sealed industrial imaging hood | Matte black enclosed box
CH3 | Optics Protection | Prevent dust/moisture degradation. | Protect optics from contamination affecting image clarity | Optics clarity status (indirect via QC checks) | AR-coated protective glass window | Acrylic protective window
CH3 | Calibration Target | Enable periodic color and scale recalibration. | Maintain grading consistency and size calibration | Color reference patch values, spatial calibration reference | X-Rite color checker + scale ruler | Printed color/gray card + scale
CH3 | Inference Engine | Output grade score, defect flags, confidence. | Convert image features into grading decisions | Grade score, defect flags, bounding boxes, confidence % | Edge AI processor (Jetson / Pi 5 + model) | MCU/Pi-based lightweight CNN
CH3 | Evidence Output | Store images + overlays + model metadata. | Ensure traceability and auditability of grading results | Image file, overlays, model version, timestamp | Signed image + metadata logging | Standard image storage with timestamp
CH4 — CHEM / Spectroscopy Screening
Channel | Functionality | Requirement | Why Do We Need It? | What Data Do We Get From It? | Industry-Grade Component | Low-Cost Component
CH4 | VIS/NIR (11-ch) | Capture compact spectral signature for proxy modeling. | Enables compact chemical proxy modeling | 11-channel spectral intensity vector | AMS AS7341 | AS7265x
CH4 | VIS Multispectral (6-ch) | Provide cost-optimized spectral classification. | Low-cost material classification | 6-band intensity values | AMS AS7262/63 | TCS34725 (limited bands)
CH4 | 18-ch Multispectral | Enable higher spectral resolution discrimination. | Improve discrimination between similar compositions | 18-band spectral vector | Dual AS7341 / AS7343 | Combined 6-ch modules
CH4 | True NIR Spectrometer | Deliver high-precision chemical fingerprinting. | High-resolution chemical composition detection | Continuous spectrum (nm vs intensity) | Hamamatsu MEMS NIR spectrometer | Entry-level MEMS NIR module
CH4 | UV Context | Detect fluorescence/surface treatment markers. | Detect whitening agents/coatings | UV response intensity / fluorescence signal | UV LED + spectral sensor | UV LED + photodiode
CH4 | Multi-Point + References | Apply dark/white correction and spatial consistency. | Correct sensor drift and ambient influence | Dark reference spectrum, white reference spectrum, corrected spectral vector | Calibrated white tile + dark shutter | Printed white card + dark capture
CH4 | Spectral Pre-Processing | Perform normalization, smoothing, T/RH compensation. | Improve classification robustness | Normalized/smoothed spectral vector | Edge processor preprocessing | MCU normalization routines
CH4 | Classification | Output composition/adulteration risk with confidence. | Convert spectra into actionable quality metrics | Composition score, adulteration index, confidence % | ML model (PLS/PCA/NN) | Rule-based classifier
CH4 | Storage Link | Sync spectra to CH7 evidence layer. | Enable traceability and dataset building | Stored spectral file + metadata | Secure data sync protocol | Standard timestamped storage
CH5 — PHYS / Process
Channel | Functionality | Requirement | Why Do We Need It? | What Data Do We Get From It? | Industry-Grade Component | Low-Cost Component
CH5 | Weight | Measure net mass and validate fill-level. | Verify fill accuracy and detect mass deviations | Weight (grams/kg), tare value | Industrial load cell + ADS1232 | HX711 + load cell
CH5 | Firmness | Quantify texture/ripeness via mechanical sensing. | Measure texture/ripeness consistency | Applied force (N), deformation response | Industrial force sensor | FSR (Force Sensitive Resistor)
CH5 | Temperature | Monitor product thermal condition. | Detect improper storage or transport temperature | Product temperature (°C) | PT100/PT1000 probe | DS18B20 digital sensor
CH5 | Vibration/Shock | Detect transport/handling stress context. | Identify mechanical stress affecting quality | Acceleration (g), shock events | Industrial IMU (Bosch BMI series) | MPU6050
CH5 | Derived Metrics | Compute density proxy and batch KPIs. | Generate higher-level quality indicators | Density estimate, batch variation metrics | Edge processor computation | MCU-based computation
CH5 | Benchmark Rules | Flag out-of-spec deviations. | Detect compliance failures | Out-of-spec flag, deviation percentage | Configurable rule engine | Fixed firmware thresholds
CH6 — PM / Dust
Channel | Functionality | Requirement | Why Do We Need It? | What Data Do We Get From It? | Industry-Grade Component | Low-Cost Component
CH6 | PM1/PM2.5/PM10 | Measure airborne particulates affecting chamber integrity. | Detect airborne contamination affecting sensors | PM1, PM2.5, PM10 concentration (µg/m³) | Sensirion SPS30 | Plantower PMS5003
CH6 | Device Contamination | Detect optics/sensor fouling risk. | Identify gradual contamination impacting accuracy | Dust trend over time, fouling indicator | SPS30 trend monitoring | PMS5003 trend logic
CH6 | Inlet Protection | Prevent dust ingress into sensing chamber. | Protect internal sensors from contamination | Filter status (via DP/PM trend) | Replaceable HEPA filter | Mesh + foam filter
CH6 | Flow Conditioning | Stabilize airflow for accurate PM measurement. | Ensure stable sampling conditions | Flow-aligned PM reading stability | Controlled airflow channel design | Basic duct + fan
CH6 | Confound Guards | Reduce VOC false positives due to dust interference. | Avoid misclassification due to particulate interference | Confound flag, adjusted VOC confidence | Fusion logic (PM + VOC correlation) | MCU rule-based adjustment
CH6 | Maintenance Intelligence | Track dust dose and filter replacement threshold. | Enable predictive maintenance | Cumulative dust exposure, replacement alert flag | Dust dose accumulation model | Runtime-based replacement counter
CH7 — Platform / Logs / Evidence
Channel | Functionality | Requirement | Why Do We Need It? | What Data Do We Get From It? | Industry-Grade Component | Low-Cost Component
CH7 | Device Identity | Maintain authenticated device registry and mapping. | Ensure secure authentication and device traceability | Device ID, public/private key, site/SKU mapping | Secure element (ATECC608A) + cloud device registry | MCU-based key storage
CH7 | Telemetry Ingestion | Collect structured payloads from CH0–CH6. | Centralize all channel outputs into unified format | Structured telemetry payload (sensor data + metadata) | Edge gateway with structured schema engine | MCU JSON packaging
CH7 | Evidence Packs | Build append-only, hash-signed export bundles. | Ensure tamper-proof traceability and export integrity | Evidence bundle (sensor data, images, spectra, hashes) | Signed log bundle with SHA256 + secure storage | Hash-based file bundling
CH7 | Audit Logs | Maintain compliance-grade traceability records. | Provide regulatory-grade traceability | User actions, run history, configuration changes | Immutable cloud audit log service | Local append-only log file
CH7 | Model Registry | Track model versions and threshold lineage. | Ensure reproducibility and model governance | Model version ID, threshold policy version | Cloud ML model registry | Local version tagging
CH7 | Drift Monitoring | Monitor long-term sensor and model drift. | Detect degradation of sensors or models | Drift metrics, baseline deviation trend | Statistical drift detection engine | Rolling baseline comparison
CH7 | Dashboards & Alerts | Provide KPIs, trends, escalation workflows. | Enable monitoring and operational decisions | KPIs, trend graphs, alert states, escalation logs | Cloud dashboard + alert engine | Local web UI + basic alerts
CH8 — Connectivity + Storage
Channel | Functionality | Requirement | Why Do We Need It? | What Data Do We Get From It? | Industry-Grade Component | Low-Cost Component
CH8 | Wireless/BLE | Enable Wi-Fi sync and BLE provisioning. | Enable device provisioning and cloud synchronization | Wi-Fi RSSI, connection status, provisioning metadata | ESP32-S3 module | ESP8266 + BLE module
CH8 | Local Storage | Support offline-first buffering (microSD/eMMC). | Store runs when offline and buffer large datasets | Stored run files, storage capacity usage | Industrial eMMC / pSLC microSD | Standard microSD card
CH8 | RTC | Provide accurate timestamping and run orchestration. | Ensure precise time synchronization | Timestamp (date/time), run start/end markers | DS3231 (TCXO RTC) | MCU internal RTC
CH8 | GNSS | Enable geo-tagged traceability. | Provide location-based traceability and site validation | Latitude, longitude, altitude, UTC time | u-blox NEO-M8 series | ATGM336H module
CH8 | Integrity Upload | Ensure exactly-once, checksum-verified sync. | Prevent duplicate or corrupted uploads | Upload status, checksum hash, sync confirmation ID | TLS-secured MQTT/HTTPS with sequence control | HTTP POST with checksum retry
CH8 | Health Telemetry | Report connectivity and storage KPIs to CH7. | Monitor operational reliability | Connectivity uptime %, storage wear %, error codes | Dedicated device health monitor | Firmware-based reporting